## Prover
Prover
Options
Prover | Gravimetric |  |  |  |  | Calibration Type
Gravimetric
Inputs | [5 empty] | Gravimetric+Master meter
Inputs | [5 empty] | Prover+Densitometer
Coverage factor (k) |  | 2 |  |  |  | Prover+Densitometer+Master meter
Prover Uncertainty Parameters
Calibration reference |  | 0.025 %
T/P corrections |  | 0.02 %
Drift |  | 0.02 %
Computation |  | 0.001 %
Calibration Data for API MPMS 4.2 Repeatability Calculation
Calibration runs |  | 3
Range of results |  | 0.02%
Random uncertainty |  | 0.0293%
Master Meter Uncertainty Parameters
T/P corrections |  | 0.02 %
Drift |  | 0.02 %
Pulse count |  | 0.01 %
Computation |  | 0.001 %
Master Meter Calibration Data for API MPMS 4.2 Repeatability Calculation
Calibration runs |  | 5
Range of results |  | 0.05%
Random uncertainty |  | 0.027%
Uncertainty Budget
Prover volume flow rate
|  |  | Symbol | Source of Uncertainty | Expanded Uncertainty (Ui) | Probability Distribution | Divisor | Standard Uncertainty (ui) | Sensitivity (ci) | ci·ui
|  |  | U(cal) | Calibration | 0.025 % | Normal | 2 | 0.0125% | 1 | 0.0125% | 1.5625e-08 | Normal
|  |  | U(rep) | Random uncertainty | 0.0293% | Normal | 2 | 0.01465% | 1 | 0.01465% | 2.1462e-08 | Rectangular
|  |  | U(corr) | T/P corrections | 0.02 % | Normal | 2 | 0.01% | 1 | 0.01% | 1e-08
|  |  | U(drift) | Drift | 0.02 % | Rectangular | 1.7321 | 0.01155% | 1 | 0.01155% | 1.3333e-08
|  |  | U(comp) | Computation | 0.001 % | Normal | 2 | 0.0005% | 1 | 0.0005% | 2.5e-11
6.0446e-08
[9 empty] | √Σ(ci·ui)² | 0.02459%
Master meter volume flow rate
|  |  | Symbol | Source of Uncertainty | Expanded Uncertainty (Ui) | Probability Distribution | Divisor | Standard Uncertainty (ui) | Sensitivity (ci) | ci·ui
|  |  | U(cal) | Calibration | 0.04917% | Normal | 2 | 0.02459% | 1 | 0.02459% | 6.0446e-08
|  |  | U(rep) | Random uncertainty | 0.027% | Normal | 2 | 0.0135% | 1 | 0.0135% | 1.8225e-08
|  |  | U(corr) | T/P corrections | 0.02 % | Normal | 2 | 0.01% | 1 | 0.01% | 1e-08
|  |  | U(drift) | Drift | 0.02 % | Rectangular | 1.7321 | 0.01155% | 1 | 0.01155% | 1.3333e-08
|  |  | U(pulse) | Pulse count | 0.01 % | Rectangular | 1.7321 | 0.005774% | 1 | 0.005774% | 3.3333e-09
|  |  | U(comp) | Computation | 0.001 % | Normal | 2 | 0.0005% | 1 | 0.0005% | 2.5e-11
1.0536e-07
[9 empty] | √Σ(ci·ui)² | 0.03246%
Results
Volume Calculated Uncertainty
|  |  | Symbol | Parameter | Standard Uncertainty | Probability Distribution | Multiplier | Expanded Uncertainty
|  |  | qv | Prover volume flow rate | 0.02459% | Normal | 2 | 0.04917%
|  |  | qs | Master meter volume flow rate | 0.03246% | Normal | 2 | 0.06492%
Inputs
Densitometer
ρ uncertainty |  | 0.044%
CTPL
Calculation module |  | 0.05%
Standard density uncertainty
|  |  | Symbol | Source of Uncertainty | Expanded Uncertainty (Ui) | Probability Distribution | Divisor | Standard Uncertainty (ui) | Sensitivity (ci) | ci·ui
|  |  | U(ρ) | densitometer uncertainty | 0.044% | Normal | 2 | 0.022% | 1 | 0.022% | 4.84e-08
|  |  | U(ctpl) | ctpl uncertainty | 0.05% | Normal | 2 | 0.025% | 1 | 0.025% | 6.25e-08
1.109e-07
[9 empty] | √Σ(ci·ui)² | 0.0333%
Prover + densitometer mass flowrate
|  |  | Symbol | Source of Uncertainty | Expanded Uncertainty (Ui) | Probability Distribution | Divisor | Standard Uncertainty (ui) | Sensitivity (ci) | ci·ui
|  |  | U(cal) | Calibration | 0.025 % | Normal | 2 | 0.0125% | 1 | 0.0125% | 1.5625e-08
|  |  | U(rep) | Random uncertainty | 0.0293% | Normal | 2 | 0.01465% | 1 | 0.01465% | 2.1462e-08
|  |  | U(corr) | T/P corrections | 0.02 % | Normal | 2 | 0.01% | 1 | 0.01% | 1e-08
|  |  | U(drift) | Drift | 0.02 % | Rectangular | 1.7321 | 0.01155% | 1 | 0.01155% | 1.3333e-08
|  |  | U(comp) | Computation | 0.001 % | Normal | 2 | 0.0005% | 1 | 0.0005% | 2.5e-11
|  |  | U(ρs) | Standard density uncertainty | 0.0666% | Normal | 2 | 0.0333% | 1 | 0.0333% | 1.109e-07
1.7135e-07
[9 empty] | √Σ(ci·ui)² | 0.04139%
Prover+densitometer+Master meter mass flow rate
|  |  | Symbol | Source of Uncertainty | Expanded Uncertainty (Ui) | Probability Distribution | Divisor | Standard Uncertainty (ui) | Sensitivity (ci) | ci·ui
|  |  | U(cal) | Calibration | 0.08279% | Normal | 2 | 0.04139% | 1 | 0.04139% | 1.7135e-07
|  |  | U(rep) | Random uncertainty | 0.027% | Normal | 2 | 0.0135% | 1 | 0.0135% | 1.8225e-08
|  |  | U(corr) | T/P corrections | 0.02 % | Normal | 2 | 0.01% | 1 | 0.01% | 1e-08
|  |  | U(drift) | Drift | 0.02 % | Rectangular | 1.7321 | 0.01155% | 1 | 0.01155% | 1.3333e-08
|  |  | U(pulse) | Pulse count | 0.01 % | Rectangular | 1.7321 | 0.005774% | 1 | 0.005774% | 3.3333e-09
|  |  | U(comp) | Computation | 0.001 % | Normal | 2 | 0.0005% | 1 | 0.0005% | 2.5e-11
2.1626e-07
[9 empty] | √Σ(ci·ui)² | 0.0465%
Gravimetric Master meter mass flow rate
|  |  | Symbol | Source of Uncertainty | Expanded Uncertainty (Ui) | Probability Distribution | Divisor | Standard Uncertainty (ui) | Sensitivity (ci) | ci·ui
|  |  | U(cal) | Calibration | 0.05% | Normal | 2 | 0.025% | 1 | 0.025% | 6.25e-08
|  |  | U(rep) | Random uncertainty | 0.027% | Normal | 2 | 0.0135% | 1 | 0.0135% | 1.8225e-08
|  |  | U(corr) | T/P corrections | 0.02 % | Normal | 2 | 0.01% | 1 | 0.01% | 1e-08
|  |  | U(drift) | Drift | 0.02 % | Rectangular | 1.7321 | 0.01155% | 1 | 0.01155% | 1.3333e-08
|  |  | U(pulse) | Pulse count | 0.01 % | Rectangular | 1.7321 | 0.005774% | 1 | 0.005774% | 3.3333e-09
|  |  | U(comp) | Computation | 0.001 % | Normal | 2 | 0.0005% | 1 | 0.0005% | 2.5e-11
1.0742e-07
[9 empty] | √Σ(ci·ui)² | 0.03277%
Results
Mass Calculated Uncertainty
|  |  | Symbol | Parameter | Standard Uncertainty | Probability Distribution | Multiplier | Expanded Uncertainty
|  |  | Qpm | Prover mass flow rate | 0.04139% | Normal | 2 | 0.08279%
|  |  | Qpmm | Master meter mass flow rate | 0.0465% | Normal | 2 | 0.09301%
|  |  | Qgmm | Gravimetric Master meter | 0.03277% | Normal | 2 | 0.06555%